The Eyegaze Edge must be disinfected prior to loaning it to a user. The following procedure will clean the system without harming it. *Please note: this procedure is intended only for cleaning the Eyegaze Edge. We cannot guarantee it will be safe to use on other systems.*

|  |  |
| --- | --- |
| **Step 1:**Disconnect the camera from the Eyegaze tablet. | **Step 2:**Wipe the tablet surface with70% or 90% isopropyl alcohol. Wipe the tablet’s front, back, and edges, paying close attention to ensuring no alcohol gets into the ports. Putting alcohol inside of a port can cause serious damage to the computer.  |
| **Step 3:**Wipe all the camera surfaces with isopropyl alcohol, being sure to avoid getting alcohol into the port and camera aperture. | **Step 4:**Wipe all cables and connectors with isopropyl alcohol.  |
| **Step 5:**Clean the mount. You may use either isopropyl alcohol or any antibacterial wipe, being sure to wipe all surfaces. | **Step 6:**Be sure all components are completely dry before packing. |